## Needs
Need ID | Description | Justification | Required
1 | The wheelchair must be safe to operate at all times. | The user will participate in high-contact, fast-paced court sports where collisions, rapid acceleration, and tipping are common. | Y
2 | The wheelchair must fail safely in the event of power or control system failure. | Loss of power or sensor failure during high-speed athletic activity could cause serious injury. | Y
3 | The wheelchair must respond quickly and intuitively to hands-free user input. | The core purpose of the design is to free the user’s hands for intuitive athletic performance. | Y
4 | The wheelchair movement must be primarily controlled by leaning. | The client explicitly requested posture-based control to eliminate hand propulsion during sport. | Y
5 | The wheelchair must be highly stable and resistant to tipping. | Stability is critical in wheelchair athletics to prevent loss of control. | Y
6 | The wheelchair must stay upright after collisions to the wheelchair body and user. | Collisions with other athletes are expected in wheelchair sports. The system must maintain upright orientation or recover quickly to prevent interruption of play. | Y
7 | The wheelchair must withstand repeated impacts without functional degradation. | Collisions are expected under normal use. | Y
8 | The wheelchair should allow limited manual control or override. | Manual control provides redundancy in case of power failure and allows repositioning in non-athletic settings where powered motion may not be appropriate. | Y
9 | The wheelchair must fit through, or be easily modified to fit through, standard doorways and common indoor spaces. | Although athletics are the primary function, basic accessibility is required. | Y
10 | The wheelchair must operate long enough to support a full athletic session. | The client expects the system to function for the duration of practices or games without interruption. | Y
11 | The wheelchair must be suitable for competitive indoor court surfaces. | Basketball and pickleball courts have traction, marking, and floor protection constraints. | Y
12 | The wheelchair must not damage regulation court surfaces. | Gym facilities are stakeholders. Damage to finished court surfaces could prevent use in competitive venues and reduce acceptance by organizations. | Y
13 | Seating must include pressure-relief features to reduce risk of pressure sores. | The user has lower-limb paralysis and may have reduced sensation and circulation. Extended seated use during athletics increases risk of pressure injury, which is a known medical concern for wheelchair users. | Y
14 | The wheelchair remains intuitive to set up and take down without the use of tools. | The user and caretaker must independently transport and store the wheelchair. | Y
15 | The wheelchair must have integrated battery power. | It is essential that the wheelchair uses a battery to move. | Y
16 | The wheelchair may allow adjustable seating or scalability for different users. | The user is young, and will grow more. Adjustability allows for use when the user is larger. | N
is "quick" a want?

## Specs
Spec ID | Need ID Ref. | Description | Justification | Required | caleb notes | charles notes
A | 1,7 | The wheelchair shall withstand repeated expected sport-contact loading without fracture, detachment of critical components, or permanent deformation that causes failure of any applicable safety, stability, or control specification. | The wheelchair will be used in a high-contact athletic environment where repeated impacts and loading are expected during normal use. Structural failure or permanent deformation could create immediate safety hazards and make the system unusable during play. | Y
B | 1,2 | Upon sensor fault, controller fault, propulsion power interruption, or triggering of emergency stop device, the wheelchair shall automatically enter a safe state within 0.10 s and reduce commanded wheel speed to less than 5% of maximum within 0.50 s. | The wheelchair must fail safely during electrical or control faults to reduce the risk of unintended motion. A rapid transition to a safe state is necessary because the device may be operating at speed in close proximity to other players. | Y
C | 1,5 | The wheelchair shall remain upright during maximum acceleration, maximum braking, and maximum turning, with no wheel lifting more than 25 mm off the ground for longer than 0.50 s. | Stability is essential for safe athletic use because tipping or excessive wheel lift can cause loss of control, user injury, or interruption of play. The 25 mm wheel-lift threshold was selected as a conservative engineering limit to identify unacceptable instability during aggressive maneuvers. | Y
D | 1,6 | The wheelchair shall remain upright and operable after expected frontal or lateral sport contact, and shall not allow disloging of seated user during such events. | Collisions with other players and wheelchairs are expected during court sports. The system must preserve upright orientation and user containment during routine contact so that normal gameplay does not create unacceptable safety risk. | Y | are there any standards here
E | 1 | The wheelchair shall have no exposed sharp edges, burrs, or accessible pinch points in normal operation, and no user-contact surface shall exceed 43°C during operation. | The user will remain in direct physical contact with the wheelchair throughout operation. Eliminating cut, pinch, and burn hazards is necessary to ensure the system is safe for repeated athletic use. [28] | Y
F | 3 | The wheelchair shall achieve an average forward acceleration of at least 2.85 m/s^2 from rest over the first 5 m of travel. | Wheelchair basketball sprint studies report a 5 m sprint time of approximately 1.87 s for competitive players, corresponding to an average acceleration of about 2.86 m/s^2 [29]. | Y | need to verify feasability with spreadsheet
G | 3 | The wheelchair shall achieve a maximum forward speed of at least 11.0 mph. | Elite wheelchair basketball sprint testing reports peak velocities of approximately 5.3 m/s (≈11.9 mph). A target of 11 mph is therefore slightly conservative relative to elite performance [30]. | Y
H | 3,4 | The wheelchair shall achieve a angular velocity of at least 200 deg/s within 0.50 s from neutral motion. | Elite wheelchair basketball sprint data show athletes reach approximately 3.4 to 4.0 m/s within the first 5 m [30], and wheelchair court-sport agility data report turning radii of approximately 0.94 to 0.98 m [31]. These values correspond to turning rates on the order of 200 to 240 deg/s, so 200 deg/s was selected as a conservative, data-driven target. | Y
I | 4 | The wheelchair shall provide all normal propulsion and steering control through torso-lean input alone during standard athletic operation, without requiring hand propulsion for commanded motion. | The primary purpose of the design is to allow the user to control the wheelchair through torso lean so that the hands remain available for athletic activity. | Y
J | 4 | The lean-control system shall be calibratable for a given user in 5 min or less. | The control system must be adjustable to the user’s available torso motion without requiring excessive setup effort before operation. | Y
K | 4 | The lean-control system shall detect and respond to intentional torso-lean input within 0.10 s of threshold crossing. | Responsive control input is necessary for intuitive and effective maneuvering during athletic operation. | Y
L | 8 | The wheelchair shall be manually movable by a caregiver when unpowered on a level indoor surface. | The wheelchair must remain transportable and repositionable when electrical power is unavailable. | Y
M | 8 | The shall be a backup method for controlled low-speed repositioning of the wheelchair at speeds of at least 1.0 mph when the primary lean-control input is unavailable or disabled. | A backup mobility function is necessary to preserve basic controlled movement during troubleshooting or primary-control failure. | Y
N | 9 | In its transport configuration, the wheelchair shall have a maximum width of 31.5 in. | The wheelchair must be capable of passing through standard accessible doorway openings during transport. [32] | Y
O | 10 | The wheelchair shall operate for at least 2.0 h on a full battery charge under normal athletic use without requiring recharge or battery replacement. | The wheelchair must support a full practice or game session without interruption so that battery limits do not prevent normal use. | Y
P | 11 | From maximum operating speed, the wheelchair shall come to a controlled stop in 3.0 s or less on a clean indoor court surface without prolonged uncontrolled skidding or loss of directional stability. | Controlled braking on indoor court surfaces is necessary to maintain user safety and predictable maneuvering during athletic operation. The 3.0 s threshold was selected as a conservative engineering target for safe stopping during sport use. | Y
Q | 12 | After 2.0 h of representative indoor court operation, the wheelchair shall cause no visible gouging, chipping, tearing, or persistent tire marking on the court surface. | Gym floors are an important stakeholder concern. The wheelchair must be usable in athletic facilities without causing unacceptable surface damage that could prevent adoption. | Y
R | 13 | The seating system shall provide pressure-management support suitable for prolonged wheelchair use and shall allow the user to perform an independent forward or lateral pressure-relief maneuver without leaving the chair. | Because the user is at elevated risk of pressure injury during prolonged sitting, the seating system must support routine pressure redistribution and compatible pressure-relief maneuvers. [33] [34] | Y
S | 14 | The wheelchair shall be capable of setup from transport state to ready-to-use state in 5 min or less without tools. | The wheelchair must be prepared for use efficiently without specialized equipment. | Y
T | 14 | The wheelchair shall be capable of takedown from ready-to-use state to transport state in 5 min or less without tools. | The wheelchair must be returned to transport condition efficiently without specialized equipment. | Y
U | 15 | All propulsion batteries, motor controllers, and required electronics shall be mounted on-board the wheelchair during normal operation, with no external trailing power cables or off-board battery pack. | The wheelchair must function as a self-contained powered system during athletic use. External power hardware would create safety, mobility, and usability issues. | Y
V | 16 | The wheelchair shall allow adjustment of seat width, footrest position, or seat-back position to accommodate users with body dimensions varying by at least 20% from the primary client baseline. | Limited adjustability improves long-term usefulness and allows the system to remain functional as the user grows or if minor fitting changes are needed. | N

## Specs_Draft (hidden)
Spec ID | Related Need ID | Specification | Justification and Reference | Req?
A | 1 | The wheelchair must not cause an increased likelihood of injuries compared to other sports wheelchairs cause 0 injuries to anyone No injuries can be caused by failure in wheelchair functionality Injuries can only be caused by operator’s misuse of device | Safety is of the highest importance to our design and increased injuries to operators or others will not be permitted. (aka Bodily harm is non-ideal) | Y
B | 2 | The center of mass of the rider must undergo no more than 9.81 m/s^2 in the event of power or control system failure | <9.81 m/s^s> is within tolerable limits and will prevent injuries even without power (aka No fall over) | Y
| 1,2 | There must be 0 injuries in the event of power or control system failure. | Bodily harm no bueno
C | 3 | The wheelchair must reach linear velocities of 9.3 mph in under 3 seconds. | By giving the capability to accelerate within these specs, the chair is able to be responsive and competitive for recreational use (aka We did da mafs) | Y
D | 3 | The wheelchair must reach an angular velocity (measured about a vertical axis through the center of mass) of 180°/s within .5 seconds | This specification will allow the maneuverability to (aka I spun myself and it seemed reasonable.) | Y
F | 4 | An operator lean of up to 85° foreward and up to 60° backward shall cause the wheelchair to translate forward or backward, at maximum speed. | These parameters allow the operator of the wheelchair to have the full range of control within a reasonable degree of leaning (Aka lean=zoomies) | Y | these need to be refiend
H | 4 | An operator lean of up to 30° left or right shall cause the wheelchair to rotate left or right, with a 30° lean corresponding to an angular velocity of 180°/s mph | These parameters allow the operator of the wheelchair to have the full range of control within a reasonable degree of leaning (aka lean=zoomies) | Y | these need to be refiend
I | 5 | Under maximum deceleration from maximum velocity, the wheelchair shall not lean more than 5° | Ensuring that the wheelchair will remain stable during the most extreme operational capabilities of the wheelchair will keep the operator from tipping during general use | Y | -
J | 1,6 | The wheelchair must not lean more than 5° from head on collisions with other wheelchairs at max velocity | This specification ensures that the wheelchair will remain stable in the most extreme impacts during recreational use | Y
K | 1,6 | The wheelchair must not lean more than 5° for impacts of max velocity from its sides | This specification ensures that the wheelchair will remain stable in the most extreme impacts during recreational use | Y
L | 1,6 | The wheelchair must not lean forward more than 5° for impacts of max velocity from behind. | This specification ensures that the wheelchair will remain stable in the most extreme impacts during recreational use | Y
M | 7 | Wheelchair must work through normal in game collisions for a minimum of 2 hours | Contact is normal in wheelchair basketball, and for this design we expect the chair to be able to work for a minimum length of 1 game with no functionality issues | Y
N | 7 | The wheelchair must retain acceleration and max velocity performance characteristics after 20 impacts. |  | Y
O | 1,8 | Wheelchair must allow for 100% control of translation and rotation via manual control | The capability of manual control functions as a redundancy to ensure that the operator will always be able to control the wheelchair, even without power | Y
P | 9 | Wheelchair must be capable of fitting through an X by X space | This size constraint will allow for the transportation of the wheelchair within a building and for vehicular transportation | Y
Q | 10 | The wheelchair must be operable for a minimum of 2 hours on one battery charge | By choosing a duration of 2 hours for the battery charge, it is ensured that the wheelchair will be able to function at peak performance through at least halftime, allowing the battery to be replaced with a fully charged battery | Y
| 11 | Wheelchair must be able to accelerate from stop without skidding on court surfaces | The design should allow for smooth movement on court surfaces without slippage during gameplay
R | 12 | Wheelchair must cause no visible damage to court surface | The design should not cause more than reasonable wear on court surfaces, as this will reduce the property damage that occurs (aka No sue be good) | Y
S | 13 | The wheelchair must have at least x amount of padding User must sustain 0 pressure related sores after two hours of operation. | This specification ensures that the operator will not develop pressure-related sources from reasonable use of the wheelchair | Y
T | 14 | Solution can be deployed within 5 minutes by a skilled technician. | This specification will ensure that the wheelchair will only require a reasonable amount of time to set up for a game | Y
U | 14 | A technician can be trained in no more than 15 minutes to deploy the wheelchair | This specification will support simplicity in the wheelchair’s removable components. | Y
| 15 | The wheelchair must have an onboard battery sufficient to operate the wheelchair for the duration of 2 hours | This battery life will give sufficient time for the wheelchair to operate between breaks in competition (aka Long live zoomies!) | Y
| 16 | Users whose butts range from x to y in can fit |  | N